REPORT OF UNDERWAY pCO2 MEASUREMENTS IN SURFACE WATERS

AND THE ATMOSPHERE DURING March-April 2007

R/V Laurence M. Gould 07/4

S. C. Sutherland, Timothy Newberger, Taro Takahashi, and Colm Sweeney

Lamont-Doherty Earth Observatory of Columbia University

Palisades, NY 10964

13 January, 2015

Part 1 - Data Processing and Quality Control.

1.  General

a)  Cruise Overview

1)  This is a routine supply Palmer Station cruise, another Palmer Shuttle.  The ship departed Punta Arenas, Chile on 23 March 2007.  She arrived at Palmer Station on 28 March and departed on 2 April.  She arrived in Punta Arenas on 6 April 2007.

2) The ship's barometric pressure showed unusual offsets and variability and was replace by a constant offset from our barometer of 3.0 mb.

3)  There is a period during day 86 where the pCO2 was dramatically higher than our high standard.  We examined the response curve during this period and are confident that extrapolating is valid.  During this period an addition 2 ppm of uncertainty should be applied for a total uncertainty of ± 4 ppm.  This ship was in the vicinity of Deception Island in relative shallow and protected water.  None of our normal quality control indicators pointed to any system problems, so we assume these values are real.

4)  This file has been reprocessed during January 2015.  The timing of the seawater temperature data measured with ThermoSalinoGraph (TSG) unit was adjusted to match the timing of input temperature (Teq) of the pCO2 equilibrator in order to account for the transit time of water from the ship’s intake to the pCO2 system.

1.	b) Data Overview

1)  The pCO2 data are very good.  See below for details.

2)  The temperature probe in the pCO2 system is normally adjusted using a set of calibration constants.  The data look fine.

3)  The atmospheric pCO2 observations were reasonably well behaved, if the relative wind direction is used as a selection criteria.  Our most recent algorithm for processing these data seem to give a reasonable subset.  The air CO2 concentration (dry air) varied from about 376.5 to a high near 383.  We continue to use the Globalview CO2 data for calculating air-sea pCO2 difference, primarily for consistency with earlier cruises.

2.  Standards:

Cylinder		Concentration (ppm)

Nitrogen (UHP)	0.0

CA06684		120.66

CA06616		359.49

CA06653		252.31

CA07220		505.10

CA06644		385.30

A 4th order calibration curve is used with 5 standards, not including the Nitrogen, as described in Section 2.  The calibration curves are also plotted in Section 4.

Standards flow for about 120  seconds before an observation is made.,

3.  Details of pCO2 data.

a)  There were 2,637 individual seawater pCO2 observations.   No records were rejected.

b)  Reinterating the comments in the introduction, the period during day 86 with the very high surface pCO2 we believe to be real.  None of our quality control indicators pointed to any system problems, and the location of the data, near and "behind" Deception Island, and in relatively shallow water convince us that the data are accurate.  Because the values are well outside our high standard, there is some additional uncertainty associated with extrapolating on the response curve.  An examination of the response curve shows it to be very well behaved, and we thing extrapolation is warranted.  An additional ± 2 ppm uncertainty in the VCO2 calculation results in a total uncertainty of ± 4 ppm during this period.

4.  Details of temperature data.

a)  The temperature of water in the equilibrator is monitored continuously using a Guildline platinum resistance thermometer (readable to 0.05 oC and regularly calibrated against a mercury thermometer traceable to N.I.S.T.) and recorded on the data-logging computer.  The calibration curve is integrated into the LabVIew program that operates the system and acquires the data.  The pCO2 temperature is then a calibrated value.  We make an effort to check this calibration during cruises against a mercury thermometer.

b)  The Documentation file for the TSG was not included in our data set.

5.  Details of Salinity data.

No Autosal analyses were performed during this cruise, so the salinity must be considered uncalibrated.  However, except for one cruise where the offset was ~0.1 units, the TSG salinity has been within ±0.05 units on the Practical Salinity Scale of the autosal determinations in the four years of this program.

6.  Details of the air data

There was substantial spiking in the air data.  It remains somewhat problematic to filter the "bad" data out of these observations.  Using the relative wind and a 45° window around dead ahead gives a resonable subset of data.  Note though, that there are many reasonable values which were rejected by this same criteria.  There are a total of 532 air observations, we reject 322, resulting in 210 records meeting the acceptance criteria.

### Part 2  Description of Instrumentation

a)  The LDEO Underway System for Surface Water pCO2 Measurements:

The underway pCO2 measurement system for surface waters consists of a) a water-air equilibrator, b) an analyzer system based on a non-dispersive infra-red CO2 gas analyzer, c) sample pumps, d) 5 calibration standards and e) a data logging system.  The measurement system is schematically shown in Fig. 1, and is similar with the one described in Bates et al. (1998).  Each of these units and the data reduction procedures used will be described below.

Figure 2:  Diagram of the underway pCO2 system used for surface water pCO2 measurements during NBP04-03 which is essentially identical to the Gould System.

b)  Water-air Equilibrator:

The equilibrator has a total volume of about 30 liters and is equipped with a standpipe that is isolated from the headspace and automatically maintains a constant water level at about half the height of the equilibrator leaving about 15 liters of headspace (Figure 2).  Seawater from the ship’s uncontaminated water line is continuously pumped into the equilibrator at a rate of about 10 liters/min, giving a mean residence time of water in the equilibrator of about 1.5 minutes. The headspace above the water serves as an equilibration chamber.  A carrier gas (commonly marine air) is continuously circulated to and from the head space in a closed loop by a diaphragm pump, and exchanges CO2 with the continuous flow of seawater sprayed into the chamber through a spiral cone spray nozzle.  Because of large gas-water contact areas created by fine water droplets as well as gas bubbles in the pool of water, CO2 equilibration between the carrier gas and seawater is achieved rapidly with a e-folding time of 2 to 3 minutes.  Under normal sampling conditions, about 50 ml/min of this sea water equilibrated air is directed into the infra-red gas analyzer.  The air drawn for analysis is replaced by a controlled leak into the equilibrator through the water manometer.  At this rate, the residence time of the carrier gas in the equilibration chamber is about 300 minutes, that is about 100 times as long as the equilibration time.  Therefore, the carrier gas in the head space should always be in equilibrium with the seawater. The over all response time of the equilibrator system has been estimated to be of an order of several minutes.  The large volume of water in the equilibrator was chosen to provide a large thermal inertia in the equilibrator, so that the effects of room temperature changes on the equilibration temperature may be minimized.  The temperature of water in the equilibrator is monitored continuously using a Guildline platinum resistance thermometer (readable to 0.05 oC and regularly calibrated against a mercury thermometer traceable to N.I.S.T.) and recorded on the data-logging computer.

The water manometer is attached to the gas intake end of the equilibrator (Figure 2) and provides a visual confirmation that marine air is being returned to the equilibration chamber at the same flow rate that it is being diverted to the CO2 analyzer.  The manometer also isolates the equilibrator head space from free exchange with the ambient marine air.  The low head of the water manometer ensures that the equilibration chamber pressure is nearly identical to ambient atmospheric pressure.  The total pressure at which the gas was equilibrated, can then be measured outside the equilibrator using a precision electronic barometer (Setra Model 270, Action, MA).  This equilibration pressure is also logged on the computer.

The temperature and salinity of seawater at the in situ conditions were measured using a SeaBird Model SBE-21 thermosalinograph aboard the N. B. Palmer and a SIO/ODF thermosalinograph unit based on Neil Brown sensors aboard the R. Revelle. The precision of the reported temperature data has been estimated to be about 0.005 oC.

c)  Infra-red CO2 Gas Analyzer:

The equilibrated gas was passed through a water trap (to collect condensate,) and, and a reverse flow naphion dryer (PermaPure flushed with pure nitrogen gas) to remove water vapor (to a level of –20oC dewpoint), and was introduced into the IR sample cell at a rate of about 50 ml/min for CO2 determinations. A LI-COR infra-red gas analyzer (Model 6251, Lincoln, NB) was used. After a 2-minute purging period, the gas flow was stopped and readings were recorded on the computer. Although the CO2 analyzer has an electronic circuit to linearize the CO2 response, it exhibited a few inflexions that deviated from linearity by a few ppm.  Therefore, the pCO2 system has been designed to use the raw voltage output of the CO2 analyzer to establish our own calibration curves using five standard gas mixtures (one pure nitrogen and four CO2-air mixtures).  The CO2 concentrations in the gas mixtures were calibrated using the SIO standards determined by C. D. Keeling’s group using the manometric method.  The concentrations of CO2 in the standard gas mixtures were summarized above.

During normal operations approximately every 75 minutes each of the standard gas mixtures was passed through the analyzer for 120 seconds at a rate of about 60 ml/min.  This flow rate and time was redundantly sufficient to completely purge the IR

analyzer cell with the new gas.  The flow was stopped for 10 seconds to allow the pressure in the cell to equilibrate to atmospheric pressure and the average of 20 separate millivolt readings from the analyzer was recorded.  Following the five standard gases, samples of seawater-equilibrated air and ambient marine air were pumped through the analyzer for 120 seconds (2 minutes) each at rates of about 50 ml/min and 60 ml/min respectively to purge the previous sample in the IR cell.  The flow was stopped for 10 seconds and the average of 20 separate millivolt readings from the analyzer was recorded.  This procedure was intended to eliminate errors due to fluctuations of the dynamic pressure within the IR cell by irregular gas flow rates.  The relatively slow flow rates used for seawater-equilibrated air samples were required for the removal of water vapor using the PermaPure membrane dryer.  We analyze 30 seawater-equilibrated air samples and 5 ambient marine air samples between two sets of calibration runs using the five standard gases.

d)  Data Logging System:

The following values were recorded on a laptop computer.  The sample locations were derived from the ships Global Positioning System (GPS).  The CO2 readings for seawater equilibrated air samples were recorded approximately every 3 minutes, and those for the standard gas mixtures approximately every 90 minutes.

Time/Date stamp (GMT),

CO2 analyzer voltage output (to 0.1 mv)

Analyzer cell temperature (to 0.1 C)

Barometric pressure (pCO2 system barometer) (to 0.1 mbar)

CO2 concentration (to 0.1 ppm, preliminary value based on the last response curve)

Equilibrator temperature (to 0.01 oC)

pCO2 (uatm, preliminary value based on the last response curve)

Gas flow rate (to 0.1 cc/min)

Selector valve position (integer)

Sample ID (standard gas cylinder numbers, seawater CO2, atmospheric CO2)

The ships RVDAS system adds several other data elements from various sensors and equipment including:  Latitude, Longitude, Thermosalinograph temperature, thermosalinograph salinity, fluorometer voltage, wind speed, wind direction, barometric pressure (at ships instrument mast), water flow in the uncontaminated seawater line.

e)  Data Reduction Procedures:

The concentration of CO2 in the sample was computed in the following manner based on the millivolt reading and time of the reading.  The millivolt reading taken for each of the five standard gases at the time of sample measurement was computed by linearly interpolating as a function of time using the readings taken before and after the respective standard gases were analyzed.  This yields an interpolated millivolt reading for each of the five standard gases at the time when the sample was analyzed.  These five values were fit to a fourth-order polynomial equation (with five constants to be determined).  This serves as the response curve.  The CO2 concentration in the sample was computed using the response curve that was established at the time of each sample analysis.  This method has been demonstrated to yield more reliable CO2 values compared with those computed, for example, using a least-squares fit of a quadratic or cubic functions to the five calibration points.  The method described above yields atmospheric CO2 values that are consistent (within 1.5 ppm) with those reported for the South Pole and the Cape Grim by the Climate Monitoring and Diagnostics Laboratory/NOAA in Boulder, CO.

The partial pressure of CO2 in seawater, (pCO2)sw, at the temperature of equilibration, Teq, in the unit of microatmospheres (atm) was computed using the expression:

(pCO2)sw @ Teq = (VCO2)eq  x (Peq – Pw), .... .........................[1]

(VCO2)eq 	=  the mole fraction concentration (ppm) of CO2 in the dried equilibrated 			      carrier gas;

Peq 		=  the barometric pressure (that is equal to the total pressure of 				    equilibration) in atmospheres; and

Pw		=  the equilibrium water vapor pressure at Teq (oC) and salinity.

The water vapor pressure was computed using the following formulation;

Pw (atm) = (1/760)x(1 – 5.368x10-4x Sal)

x EXP{[0.0039476 – (1/TK)]/1.8752x10-4}, ....... [2]

where Sal is salinity on the Practical Salinity Scale and measured using the ship’s thermosalinograph, and TK is the temperature of equilibration in oK.

The (pCO2)sw at the in situ temperature, T in situ, was computed using a constant value of 0.0423 % per oC for the effect of temperature (Takahashi et al., 1993):

(pCO2)sw @ Tin situ  = (pCO2)sw @ Teq x EXP[0.0423 x (Tin-situ – Teq)].

The value for Tin situ is taken to be the seawater temperature measured by the ship’s thermosalinograph at the time of pCO2 measurements. Teq is generally warmer than

Tin-situ by 0.5 ~ 0.8 oC.  Hence the temperature correction is normally less than 3% of pCO2 values.

The over all precision of the reported (pCO2)sw values has been estimated to be about +1.5 uatm.

f)  Atmospheric Measurements:

The air measurement system is also shown schematically in Fig. 1. Uncontaminated marine air samples were collected about 10 m above the sea surface using a 1/4” i.d. thin-walled aluminum tubing protected by an exterior plastic casing and an internal plastic film lining (Synflex, Calco Inc., PA).  The air intake was located at the middle of the foremast about 10 m above the sea surface.  A KNF Neuberger air pump located near the IR analyzer was used to pump air through the tubing and into the IR analyzer.  Even when air samples were not analyzed, air was continually flowing through the sampling line to continuously purge the long Synflex tubing with fresh marine air.  For analysis, the air sample was passed through a 0.5 micron filter to remove aerosols and a drying column to remove water vapor (the same PermaPure column as used for the seawater-equilibrated air) and introduced into the IR cell for CO2 analysis at a rate of about 60 ml/min.  For each sample the cell was purged for 120 seconds, the flow was stopped for 10 seconds and the average of 20 IR millivolt output readings was recorded on the computer.

g)  Atmospheric Data Processing:

The partial pressure of CO2 in the air, (pCO2)air, was computed in the unit of microatmospheres (atm) in the same way as that for seawater using Eq. [3] below:

(pCO2)air  = (Vco2)air  x  (Pb – Pw), .... .........................[3]

(Vco2)air 	=  the mole fraction concentration (ppm) of CO2 in the dried air sample;

Pb 		=  the barometric pressure at sea surface in atmospheres; and

Pw		=  the equilibrium water vapor pressure at Tin situ (oC) and salinity given 		                 by Eq. [2].

The precision of the atmospheric pCO2 values have been estimated to be about

+ 1 atm.

Part 3 -  References

Bates, N. R., Takahashi, T., Chipman, D. W. and Knapp, A. H.  (1998). Variability of pCO2 on diel to seasonal time scales in the Sargasso Sea. Jour. Geophys. Res., 103, 15567-15585.

CO2 Group, Lamont-Doherty Earth Observatory. (2004) “Underway pCO2  System Users Manual”,  LDEO of Columbia University, Palisade, NY, pp.23.

Takahashi, T., Olafsson, J., Goddard, J., Chipman, D. W. and Sutherland, S. C., (1993).

Seasonal variation of CO2 and nutrients in the high-latitude surface oceans: A comparative study.  Global Biogeochemical Cycles, 7, 843-878.

PART 4 - Graphics

List of Figures:

Maps showing the location of each seawater pCO2 and SST value with the data point assigned a color based on an indicated color palette.

3-panel time series of verious parameters, five each for the southbound and northbound legs.  Northbound has been deleted from the final file.

A time series showing salinity data with editing.

Part 5 - Program Listings

These listings provide detailed documentation of processes applied to the data, including the calculation of a reponse curve, editing, and certain quality control operations.